## Slide 1
Conceptual design of SoLID-TOF and aging test
Outline:
  Conceptual design of SoLID-TOF
  Discussion of MRPC intrinsic resolution
  Aging test of glass
Aging test of MRPC
Conclusions
  Wang Yi
Department of Engineering Physics, Tsinghua University

## Slide 2
Motivation
JLab 6 GeV: precision measurements
high luminosity (1039) but small acceptance (HRS/HMS: < 10 msr)
or large acceptance but low luminosity (CLAS6: 1034)
JLab 12 GeV upgrade opens up a window of opportunities (DIS, SIDIS, Deep Exclusive Processes) to study valence quark (3-d) structure of the nucleon and other high impact physics (PVDIS, J/y, …)
High precision in multi-dimension or rare processes requires very high statistics  large acceptance and high luminosity
CLAS12: luminosity upgrade (one order of magnitude) to 1035
To fully exploit the potential of 12 GeV, taking advantage of the latest technical (detectors, DAQ, simulations, …) development
SoLID: large acceptance detector can handle 1037 luminosity (no baffles)
1039 with baffles

## Slide 3
Overview of SoLID
Full exploitation of JLab 12 GeV Upgrade
 A Large Acceptance Detector AND Can Handle High Luminosity (1037-1039)
Take advantage of latest development in detectors , data acquisitions and simulations
Reach ultimate precision for SIDIS (TMDs), PVDIS in high-x region and threshold J/y
5 highly rated experiments approved (+3)
Three SIDIS experiments, one PVDIS, one J/y production (+ three run group experiments)
Strong collaboration (250+ collaborators from 70+ institutes, 13 countries)
Significant international contributions (Chinese collaboration)
Solenoidal Large Intensity Device

## Slide 4
Main requirements for TOF
The MRPC is developed for the TOF of SoLID
Main Requirements for TOF:
  /k separation up to 2.5GeV/c
  Time resolution < 100ps
  Rate capability > 10kHz/cm²
To separate very few K from , we require the TOF time resolution as good as possible.
We have to improve the time resolution to about 20ps.

## Slide 5
Conceptual design of SOLID-TOF
MRPC TOF wall we designed contain 150 MRPC modules in total, with 50 gas boxes and 3 counters in each box, covering the area of 10m².
MRPC1
MRPC2
MRPC3
100cm
16cmm
28cm
Total channel ~3600

## Slide 6
6
Introduction of MRPC
Application in nuclear physics experiments
Application in industry（Muon tomography)
Application in medicine（TOF-PET）
6
Standard parameters:
Resistivity of glass: ~1012 Ω.cm
Time resolution <100ps
Efficiency >95%
Dark current: a few nA
Noise <1Hz/cm2
MRPC is made of thin
glass, large area and cheap
The inner glasses are
floating, take and keep
correct voltage by
electrostatics. it is
transparent to fast signals
Thin gap->good timing
Multi-gap-> high efficiency
Read out
Read out
Insulator
Carbon
Insulator
Resistive plate

## Slide 7
Volume resistivity: ~1010.cm
363mm
219mm
171mm
25mm
 | Length/mm | Width/mm | Thickness/mm
Gas gap | - | - | 0.25×10
Inner glass | 320 | 130-171 | 0.7
Outer glass | 330 | 138-182 | 1.1
Mylar | 335 | 153-198 | 0.18
Inner PCB | 350 | 182-228 | 1.6
Outer PCB | 350 | 172-218 | 0.8
Honeycomb | 330 | 153-198 | 6
Material dimensions
A MRPC prototype for SoLID-TOF

## Slide 8
Cosmic test
Cosmic ray test@Tsinghua August, 2011
Efficiency > 95% @ 96kV/cm (6.0kV)
Time resolution: ~ 52ps
PMT3
PMT4
PMT0
PMT2
PMT1
MRPC
Cosmic Ray Muon

## Slide 9
Beam test @ Hall A
Shield
High Energy e-
Top View
PMT1
PMT2
PMT3
PMT4
PMT0
Target
Target
Test setup
The diagram of DAQ system

## Slide 10
Rate Performance
Voltage: 6800V
Flux:11kHz/cm²
Efficiency:~95%
Time Resolution: 78ps
Flux:16kHz/cm²
Efficiency:~95%
Time Resolution: 82ps

## Slide 11
Charge spectrum at different rate
Charge Distribution
Base Line
Signal
Significant reduction of charge distribution can be seen as the background rate increases.
The reduction of electrical field is caused mainly by both background irradiation and electron beams.
Rate capability of the counter is higher than 16kHz/cm².

## Slide 12
TOF electronics
Fast preamplifier：
Ninos (TOT) (ALICE)
Padi (TOT) (GSI)
CAD (TOT) (Tsinghua)
Low time jitter faster FEE
TDC
HPTDC (ALICE, 25ps/ch)
GET4 (GSI, 25ps/ch)
FPGA TDC (GSI and other)
Pulse wave form digitizer
DAQ

## Slide 13
Intrinsic resolution of 10gap-MRPC
Time jitter of HPTDC 25ps
NINOs ASIC+Interface card 24ps
MRPC resolution tested 40ps
So the intrinsic resolution of MRPC is:

## Slide 14
14
Resolution of 24 x140m MRPC
Fast wave form digitizer 30-> 5ps
Low time jitter fast FEE 24-> 5ps
MRPC intrinsic 20-> 13ps
Time resolution of TOF:

## Slide 15
TOF PID at the EIC
TOF Positives
Compact (10-15 cm of space needed), allowing more room for, e.g., RICH
More clean PID
Can contribute to e/h identification using TPC dE/dx
TOF Negatives
Start counter?
@ 1m distance | Distance | Pion-Kaon Separation | Kaon-Proton Separation
σtot=100 ps | 1m |  | 
σtot=10 ps | 1m |  | 
σtot=10 ps | 3m |  | 
Momentum reach of TOF PID reaches interesting levels if one can achieve ~ 10 ps

## Slide 16
16
 | MRPC ( C. Williams et al. ) | MRPC ( UIUC & BNL)
Gas Gap Width | 160um (fishing line) | 105um (fishing line, and 165, 125, 75um)
# of Gas Gaps | 4 stack x 6 layers = 24 | 4 stack x 9 layers = 36
# of thin glass layers | 4 stacks x 5 layers = 20 (250um thick glass) | 4 stack x 8 layers = 32 (210um thick glass)
Preamplifier | Differential type, NINO chip | LMH6881 2.8-GHz + Evaluation Board
TDC and DAQ | Oscilloscope (10Gs) | DRS4-V5(5Gs) + PC
Time resolution | 20 ps published (16 ps private comm) | 18ps
Configuration of eRD10 Prototype mRPC
Gas gap width : 160um
(16ps timing resolution)
Gas gap width : 105um
Crispin Williams
BNL&UIUC Prototype
4 x
4 x
PCB
glass
Gas gap

## Slide 17
17
Time Resolution Measurement at UIUC
MRPC-A
MRPC-B
DRS4-V5
(Wave form digitizer)
PC
(DAQ)
DC0,1,2
SC1
SC4
preamp
preamp
Gas Cylinder
USB
50 Ohm cable (<1m)
ch2
ch2
Muon trigger
Gas Mixture : “Freon” (R134a) 90% : Iso 5% : SF6 5%
SC2
SC3
DC3,4
Pb Block

## Slide 18
18
Difference between time of top and bottom mRPC, Δt = t2 – t1= 25 ps
  σt = Δt/√2 = 18 ps
  Used a “Leading Edge Discriminator technique” in software from waveform, slew corrected
  Difference in pathlength taken into account
  Wings possibly due to noise, cross-talk, streamers, or something else?
We believe best published result is 20 ps
  Nucl.Instrum.Meth. A594 (2008) 39-43
  Nucl.Instrum.Meth. A629 (2011) 106-110
Time resolution

## Slide 19
19
15Ghz wave form digitizer

## Slide 20
Aging test of low resistive glass
This glass was applied with 1000V for about 32days, integrated charge: 1 C/cm2
--roughly corresponding to the CBM life-time over 5 years operation at the maximum particle rate.
Glass specifications:

## Slide 21
21
21
Comparasion with other material
Morales talk @RPC2012

## Slide 22
22
Cyclotron U-120M (Řež)
Proton energy 36 MeV
Neutron production target Be
Neutron flux 108 - 1010 n/cm2/s
Neutron energy spectra 1 – 36 MeV
Neutron Test
target support
Special thanks to A. Kugler and
O. Svoboda for their strong support.
Neutron energy spectrum

## Slide 23
Test result

## Slide 24
Hign rate MRPC
X-ray Generator
130cm
60cm
Online irradiation test
This is online test system. The efficiency and time resolution can be obtained by cosmic ray while irradiated by X-rays. 0.1C/cm2 charge is accumulated in 35 days.

## Slide 25
Current，counting rate and efficiency change with X ray strength
Determination of X ray strength

## Slide 26
Performance Vs Dosage
By monitoring the performance on line, we can grasp the detector condition simultaneously.

## Slide 27
Performance Vs Dosage
~500 events/5days
More events ->Better statistics

## Slide 28
Charge spectrum
With the increase of dosage, the average charge decreased gradually, this is mainly from space charge effect.

## Slide 29
Conclusions
A SoLID-TOF MRPC prototype is developed, its time resolution is 75ps, rate capability reach 16kHz/cm2.
MRPC intrinsic resolution can reach 13ps. Fast FEE and high band width wave form digitizer has to be used to improve TOF resolution -> 15ps.
From high voltage and neutron radiation test, we did not find any aging problem for low resistive glass.
We also studied the long term stability of the MRPC. 0.1C/cm2 charge is accumulated. No obvious performance degradation is observed.

## Slide 30
Thanks for your attention !